## Slide 1
SMART INDIA HACKATHON 2025
EDGE-AI
PREDICTIVE COLD CHAIN MONITOR
FreeRTOS-enabled hyper-local monitoring with on-device predictive ML
ESP32
SHT30
Decision Tree
FreeRTOS
ThingSpeak HTTPS
SENSE → PREDICT → ACT
SHT30
ESP32
ML
early warning + telemetry
SIH26215 · Theme: MedTech · Hardware · Team 14
IMPACT INNOVATORS · RV UNIVERSITY
01

## Slide 2
01 · WHY THIS MATTERS
Cold-chain failures are often detected after the damage starts
The project shifts monitoring from reactive thresholds to local prediction and resilient telemetry.
THE PROBLEM
Temperature-sensitive healthcare products can leave the 2–8°C band during last-mile movement.
A threshold alarm tells the operator that a breach is happening — not how soon it will happen.
Connectivity interruptions can break continuous remote visibility.
Irregular sampling cadence can make a trained time-series model misleading if deployed naively.
A low-cost field device needs local computation, buffering and recovery — not cloud-only intelligence.
DESIGN SHIFT
Reactive monitoring
→
Predictive edge monitoring
SENSE
2-second local sensing
ANALYZE
rules + thermal models
PREDICT
Decision Tree on ESP32
BUFFER
LittleFS during outages
SYNC
ThingSpeak REST/HTTPS
IMPACT INNOVATORS · SIH26215
02

## Slide 3
02 · SYSTEM DESIGN
One device, two cadences, one resilient data path
Fast safety logic runs continuously; the trained ML model runs on the cadence it was trained for.
SENSE
SHT30 + GPS
FAST CHECK
2 s rules
ML WINDOW
10 min buckets
LOCAL STATE
alert + risk
CLOUD
ThingSpeak REST/HTTPS
IMPORTANT: 2 s sensing ≠ 10 min ML inference
FREE_RTOS TASKING
SensorTask · Core 0
SHT30/GPS → feature windows → queue
QUEUE
CloudTask · Core 1
ThingSpeak pacing + LittleFS backlog
Mutex
sync metadata
IMPACT INNOVATORS · SIH26215
03

## Slide 4
03 · HARDWARE & FIRMWARE
Built from low-cost, commonly available modules
The MVP uses a practical sensor + MCU + cloud stack while keeping intelligence at the edge.
ESP32
MCU + Wi-Fi FreeRTOS runtime
SHT30
Temp + humidity I²C / 2 s reads
NEO-6M
GPS context UART / freshness gate
LittleFS
Persistent offline telemetry backlog
WIRING
SDA
GPIO 21
SCL
GPIO 22
GPS RX
GPIO 16
GPS TX
GPIO 17
FIRMWARE STACK
Sensor task
sampling + feature window
Decision layer
heat/freeze latch + thermal models
ML layer
embedded Decision Tree
Cloud task
15 s ThingSpeak + offline backlog
IMPACT INNOVATORS · SIH26215
04

## Slide 5
04 · WHERE THE AI IS
The learned model runs on the ESP32 — not in the cloud
The cloud is telemetry. The trained Decision Tree is embedded into firmware for local inference.
TRAIN OFFLINE
Historical thermal data
EXPORT
C/C++ tree header
ADAPT
10-min bucket cadence
INFER
ESP32 Decision Tree
WARN
2 positive windows
10 FEATURES
• temperature
• 5-sample mean
• 5-sample std
• 5-sample slope
• 10-min delta
• 20-min delta
• rising fraction
• distance to 8°C
• count above 6°C
• slope acceleration
MODEL CONTRACT
Threshold
0.85
Persistence
2 windows
Domain
current bucket < 8°C
Minimum bucket coverage: 80% · five valid buckets before first inference
IMPACT INNOVATORS · SIH26215
05

## Slide 6
05 · HYBRID EDGE ANALYTICS
ML is only one layer — the system combines learned, deterministic and thermal reasoning
Each layer answers a different operational question.
FAST RULES
I
Is the product already outside the safe band?
2 s sampling
heat/freeze latch
persistence filter
EDGE ML
I
Is the trajectory likely to cross 8°C soon?
10 min windows
Decision Tree
local predictive alert
THERMAL MODELS
H
How quickly is the condition changing / accumulating exposure?
Newton cooling
Arrhenius exposure
risk score
Operational principle: keep critical detection local so connectivity is not on the critical path for prediction.
IMPACT INNOVATORS · SIH26215
06

## Slide 7
06 · MODEL VALIDATION
Primary frozen holdout: reproducible evidence, not field-accuracy marketing
Historical Nepal thermal data were split by source file so entire files were held out from model development.
66.3%
Precision
How often a positive warning was correct
68.1%
Recall
Share of target positives detected
67.2%
F1
Balance of precision and recall
0.907
ROC-AUC
Ranking quality across thresholds
HOLDOUT
6,110
rows 12 source files
FROZEN CONFUSION MATRIX
Predicted +
Predicted −
Actual +
Actual −
175
TP
82
FN
89
FP
5,764
TN
Threshold: 0.85
WHAT THE NUMBERS MEAN
Validation source: historical Nepal thermal data.
Source-file holdout: 48 development files + 12 completely held-out files.
These metrics do not represent Bengaluru field accuracy or the final physical ThingSpeak/ML runtime.
IMPACT INNOVATORS · SIH26215
07

## Slide 8
07 · RELIABILITY & SECURITY
Designed for real field constraints: connectivity loss, cadence gaps and secure telemetry
The final software release hardens the path between sensing and cloud storage.
Offline resilience
LittleFS backlog rate-limited replay
Secure cloud path
HTTPS + DigiCert root CA
Cadence quality
80% coverage 5 valid buckets
False-alarm control
2-window persistence
Sensor recovery
re-init after 5 failures
GPS freshness
120 s age gate
Important limitation: offline buffering protects telemetry data, but not real-time remote alerting while the network is disconnected.
IMPACT INNOVATORS · SIH26215
08

## Slide 9
08 · CURRENT IMPLEMENTATION STATUS
What is verified, what is software-ready, and what still needs the physical bench
This slide is intentionally explicit about evidence boundaries.
ESP32 + SHT30 sensing
VERIFIED
Hardware
Wi-Fi path
VERIFIED
Hardware
Final secure v8.3.1 firmware
READY
Static audit
ThingSpeak physical publish
PENDING
Bench test
Physical ML runtime in 2–8°C
PENDING
~50–60 min
GPS satellite fix
DEFERRED
Time-boxed
NEXT BENCH CHECK
1
Create ThingSpeak channel + enter credentials
2
Upload v8.3.1 and verify a positive entry ID
3
Run in 2–8°C for ~50–60 min to observe physical ML inference
4
Capture ml_data_quality, probability, prediction + model version
NOT CLAIMED
Bengaluru field accuracy · battery-life measurement · remote alerts during disconnection
IMPACT INNOVATORS · SIH26215
09
v8.3.1 hardening: first-bucket guard · post-attempt cloud pacing · stream-safe LittleFS backlog

## Slide 10
09 · IMPACT & ROADMAP
From validated software to a field-ready cold-chain node
The next step is not more complexity — it is physical validation in a controlled 2–8°C environment.
IMPACT
Healthcare
Earlier intervention before a breach; reduced exposure risk.
Operations
Central telemetry + local prediction + traceable backlog.
Economics
Low-cost edge hardware designed for scalable fleet use.
ROADMAP
1
Software/model holdout
COMPLETE
2
ESP32 + SHT30 integration
COMPLETE
3
Secure ThingSpeak build
READY
4
Physical ML runtime 2–8°C
NEXT
5
Field thermal validation + GPS
LATER
SENSE → ANALYZE → PREDICT → BUFFER → SYNC
EDGE-AI COLD CHAIN MONITOR
Sources: final v8.3 firmware · frozen ML holdout evidence · SIH execution status · ThingSpeak secure setup
IMPACT INNOVATORS · SIH26215
10